Conducted Immunity Test Report

Contents

[1. Test lab facility 2](#_Toc512427982)

[Test Site: 2](#_Toc512427983)

[Facility name: 2](#_Toc512427984)

[Facility address: 2](#_Toc512427985)

[Facility site description: 2](#_Toc512427986)

[Test Standards 2](#_Toc512427987)

[EMC Directive 2014/30/EU 2](#_Toc512427988)

[Test Equipment: 2](#_Toc512427989)

[Environmental conditions: 3](#_Toc512427990)

[Test Operator and Date: 3](#_Toc512427991)

[2. Product Information 3](#_Toc512427992)

[3. EUT setup 3](#_Toc512427993)

[Photograph of EUT: 4](#_Toc512427994)

[4. Test Result 5](#_Toc512427995)

[5. Summary 11](#_Toc512427996)

# Test lab facility

## Test Site:

Facility name: Keysight Technology, Technology Order Fulfillment, Colorado Springs Hardware Test Center.

Facility address: 1900 Garden of the Gods Rd, Colorado Springs, CO. 80907

Facility site description:

The Keysight Technologies Technology Order Fulfillment Colorado Spring Hardware Test Center is a certified radiated interference testing facility, which is comply with the standard requirements defined by IEC 61326.

## Test Standards

### EMC Directive 2014/30/EU

IEC 61326-1:2012 / EN 61326-1:2013 (Basic)

Radiated Immunity Reference Standards:

IEC 61000-4-3:2016 Group 1 Class A.

The products were tested in a typical configuration with Keysight Technologies test systems. This product is intended for use in a basic electromagnetic environment.

## Test Equipment:

|  |  |
| --- | --- |
| Power Sensor Address (Low Band) | USB0::0x0957::0x2D18::MY47400241::0::INSTR |
| Power Sensor Address(High Band) | USB0::0x0957::0x2D18::MY55500012::0::INSTR |
| Signal Generator Address | GPIB0::19::INSTR |
| RF Switch Com Port | COM6 |
| Power Amplifier Address(High Band) | GPIB0::2::INSTR |
| Turn Table Address | GPIB0::8::INSTR |
| Antenna Tower Address | GPIB0::9::INSTR |
| Field Probe Address (optional) | COM7 |
| Calibration File | ./configuration/RI\_Coefficients\_300pints\_newbox\_newcable.mat |
| USB power meter 1 Lower Band | Model: U2004A Calibration date: 8/5/2017 |
| USB power meter 2 Higher Band | Model: U2004A Calibration date: 8/5/2017 |
| Signal Generator SN | Model: N5181A Calibration Data12/1/2017 |
| Test Standard | IEC 61000-4-6 / EN 61000-4-6 Basic 80MHz-6GHz: 3V/M |
| Test Site | Colorado Springs TOF Hardware Test Center |
| Test Voltage | 110V |
| Test Software Version | V1.5 |

## Environmental conditions:

Temperature: 72°F; Humidity :50 RH

Note: There will be no effect to the result due to changes in mains voltage or frequency.

## Test Operator and Date:


Operator: Clifford; Report generated at: Apr.25,2018 1:57:29 PM

# Product Information

|  |  |
| --- | --- |
| Product Name: | GMAX |
| Product Model: | UXR1004A |
| Product SN: | LP3-1 |
| Project Stage: | LP3 |
| Deliver date of the test samples: | 11/15//2017 |
| DUT power: | 208V |
| Auxiliary equipment list: | Cables USB keboard |
| Model numbers covered by the test: | N/a |
| .-Hardware Difference | N/a |
| .-Software/firmware difference: | N/a |
| .-Cosmetic difference: | NA |
| Test Configuration | maximum load condition. GMAX Sr. Frame 4 channel configuration |
| Test Software Version | V1.5 |

# EUT setup

Radiated Immunity Test Setup- Table-Top and Control Room



## Photograph of EUT:



# Test Result

**Test Parameters:**

Frequency 80 MHz - 6 GHz; 80% Amplitude Modulated at 1 KHz


Sweep Rate 2.5 Seconds dwell per 1% step, 1.73E-03 Decades per Second


Filed Strength 3 v/m


Polarity Horizontal and Vertical

Rotation Front (0°) Left (90° ) Rear (180°) and Right (270°)

|  |  |
| --- | --- |
| Frequency (Hz) | Mark |
| 80e6 | A. Pass -Normal performance within specified limits |
| 6e9 | A. Pass -Normal performance within specified limits |







|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Criteria:** | | A. Pass - Normal performance within specified limits. | | |
|  | | B. Pass - Temporary degradation, self recoverable. | | |
|  | | C. Pass - Temporary degradation, requires operator intervention. | | |
|  | | D. Fail - Not recoverable, component damage. | | |
|  | |  | | |
|  |  | |  |  |

|  |  |  |  |
| --- | --- | --- | --- |
| Field Uniformity Calibration Data | | | |
|  |  | | |
| Field Uniformity Exceptions: (>6dB) | | | |
|  |  |  |  |
|  | Horizontal Axis | 2107.0 MHz | 8.80 dB |
|  | Horizontal Axis | 5040.0 MHz | 7.13 dB |
|  | Horizontal Axis | 5499.0 MHz | 6.37 dB |
|  | Vertical Axis | None | N/A |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |

# Summary

The unit passed the Radiated Immunity Emission Test!