Conducted Emission Test Report

Contents

[1. Test lab facility 2](#_Toc450134220)

[Test Site: 2](#_Toc450134221)

[Facility name: 2](#_Toc450134222)

[Facility address: 2](#_Toc450134223)

[Facility site description: 2](#_Toc450134224)

[Test Standards 2](#_Toc450134225)

[EMC Directive 2014/30/EU 2](#_Toc450134226)

[Test Equipment: 2](#_Toc450134227)

[Environmental conditions: 2](#_Toc450134228)

[Test Operator and Date: 2](#_Toc450134229)

[2. Product Information 2](#_Toc450134230)

[3. EUT setup 2](#_Toc450134231)

[Photograph of EUT: 2](#_Toc450134232)

[4. Test Result 2](#_Toc450134233)

[5. Summary 2](#_Toc450134234)

# Test lab facility

## Test Site:

Facility name: Keysight Technology, Technology Order Fulfillment, Colorado Springs Hardware Test Center.

Facility address: 1900 Garden of the Gods Rd, Colorado Springs, CO. 80920

Facility site description:

The Keysight Technologies Technology Order Fulfillment Colorado Spring Hardware Test Center is a certified radiated interference testing facility, which is comply with the standard requirements defined by IEC 61326.

## Test Standards

### EMC Directive 2014/30/EU

IEC 61326-1:2012 / EN 61326-1:2013 (Basic) Radiated Emission Reference Standards:

CISPR 11:2009+A1:2010 / EN 55011:2009+A1:2010 Group 1 Class A. The products were tested in a typical configuration with Keysight Technologies test systems. This product is intended for use in a basic electromagnetic environment.

## Test Equipment:

|  |  |
| --- | --- |
| MXE Receiver address | USB0::0x0957::0x0f0b::MY52130013::0::INSTR |
| LISN Model | EMCO 4825/2 |
| LISN SN | SN:00130174 |
| Power Limiter Calibration | 5/1/2016 |
| Test Standard | CISPER 11 Group 1 Class A RE |
| Test Site | Colorado Springs TOF Hardware Test Center 3 Meter anechoic chamber |
| Test Voltage | 110V |

## Environmental conditions:

Temperature: 72°F; Humidity :50 RH

Note: There will be no effect to the result due to changes in mains voltage or frequency.

## Test Operator and Date:

Operator: Clifford; Report generated at: May.17,2016 8:59:04 AM


# Product Information


|  |  |
| --- | --- |
| Product Name: | Ursa |
| Product Model: | 16862A |
| Product SN: | 105 |
| Project Stage: | LP1/PP |
| Deliver date of the test samples: | 5/9/2016 |
| DUT power: | Main |
| Auxiliary equipment list: | HP Keyboard and Mouse |
| Model numbers covered by the test: | 16864A, 16862A |
| .-Hardware Difference | 4 channel and 2 channel |
| .-Software/firmware difference: | Same |
| .-Cosmetic difference: | Label is different |
| Test Configuration | Test configuration is acquizaiton mode waiting for trigger |

# EUT setup

## Photograph of EUT:

# Test Result





# Summary

The unit Passed the Conducted Emission Test in Line wire!

The unit passed the Conducted Emission Test in Neutrual wire!